Specification for Lab Suspended Solids Meter for 500-20,000 ppm Range

Application

Immediate measurement of suspended solids samples in the 500-20,000 ppm (or mg/l) range from aeration basins, RAS or influent in municipal wastewater treatment plants.  Suspended solids concentrations are measured as a function of the ability of suspended materials to absorb and reflect NIR-light (Near Infra-Red).

Sensor

The measuring principle is based on transmission of a single NIR (Near Infra-Red) light beam between two glass prisms and the ability of particles to absorb this light.  NIR light is not sensitive to color and does not contribute to biological growth as with visible light.  Sensor based on “reflection” of normal light by particles are not acceptable.  The design of the sensor should have a “notched” head design to induce velocity across the glass prism lens which  produces a self-cleaning effect and orientates the particles to flow in the same direction for better measurement.  The sensor head should be made of silicon rubber.  The sensor should be supplied with 33’ (10 m) of 9 conductor shielded cable with polyurethane jacket as standard.  The sensor should have a minimum measuring range of 0-1,000 ppm and maximum of 0-20,000 ppm.  A standard sensor must be capable of measuring full range of 0-20,000 ppm or any variation between.  For wide calibration ranges, the unit must able to accept up to five calibration points.  Any combination of these samples may be calibrated and the microprocessor will calculate the correlation factor between values.  Unit should not require recalibration more than once very 12 months due to built-in temperature and light emitting diode compensation loops.

Transmitter - Control Box

The transmitter should be microprocessor based with self-instructing menu, digital circuitry, and illuminated 16 digit LCD.  Alphanumeric window should continuously display suspended solids concentration in ppm or mg/l.  Lights on the front panel should indicate transmitting pulse, high/low alarms, and alarm relay position.  The transmitter should produce a 4-20 mA, 700 ohm galvanically isolated, output signal proportional to current suspended solids concentrations and have relay signals for high and low solids alarms.  To avoid fluctuating measuring signals, the transmitter should have dampening function adjustable from 1-30 seconds.   The enclosure should be water-tight molded polyurethane box, NEMA 4 (IP65).

The power supply must be equipped with filter, fuse and varistors for protection against power surges. The alarm contacts are normally open (dry contact), when activated (closed contact) they should have max. resistive load of 250V/2A.  The transmitter should have EEPROM memory.  Following a power failure the transmitter should start up and resume measuring without requiring recalibration.  Power supply is to be 110/1/60.

Illuminated 16 digit LCD display provides current suspended solids concentrations in ppm or mg/l, and temperature of the transmitter (approximate liquid temperature) and service mode,  should be accessible by using the up arrow key.  Alarms for high and low suspended solids concentrations must be displayed on front panel and should have active relay outputs.

Programming Module

All programming and settings are performed from the outside of the transmitter by using a self-instructing menu controlled by just four touch pad keys.  Zero point adjustment should be done using clean, deaerated water.  The control box should allow for calibration against suspended solids solutions by calibrating to one point or up to five points for wide calibration range.  Any of these five points may be entered and the unit shall calculate the correlation between the points.  The values for the calibration points which are obtained from laboratory analysis, should be able to be entered anytime after calibration since the meter correlates these to light transmission values.  Tamperproof programming feature is required to keep settings from being changed, except by authorized personnel.  In the case of power loss, an EEPROM memory should save programming during power outages.

Self-Diagnostics

The software should be of Multi-task design.  It should also contain a watch-dog function connected to the microprocessor.  The software should inform the “watch-dog” at least once per second that the device is working properly.  If it does not, then the “watch-dog” shall restart the processor in order for the unit to resume measuring.

Mounting

The transmitter is to be supplied with 12’ power cord and plastic stand for mounting control box on lab counter.

Warranty

The manufacturer shall warrant the equipment to be free of defects from workmanship and material for a period of one (1) year after shipment.

Supplier

Specification is  written  around  equipment  manufactured  by  Cerlic Controls, AB,            Model - SSM-(P/IT600,  P/N 24020.  Phone  404-256-3097 or Fax  404-256-3094

File: SPEC/SSMLAB.DOC                                                                                        9/00  Rev. 3