Specification for pH/Redox Meter with Flush

Application

Continuous on-line measurement of pH or Redox for mounting in open channels, basins or tanks, with optional flushing system.  Measuring ranges of 0(14 for pH or -2000...+2000 mV for Redox potential for municipal and industrial wastewater treatment plants.

Electrode

The principle of measurement to be based on the measurement of pH and Redox as a function of millivolt.  pH electrodes should be non-refillable gel type made of glass containing measuring and reference electrodes, which in combination provide a voltage linear to the pH.  Electrode to have male type BNC connector for attachment to amplifier.  Refillable  electrodes are not acceptable.

Amplifier

The electrode should deliver a voltage, which is proportional to the pH or Redox value.  An amplifier is to be supplied to convert the pH & redox signals to RS485 output signals to a control box.  Amplifier to be supplied with 33’ (10 m) cable and M12 connector for attachment to the control box.  Cable can be lengthened with extension cables up to 165’ (50m).  Amplifier to have female BNC connector for attachment to pH or Redox electrodes.   Amplifier to be mounted in 39” long SS housing with optional flushing system.  Flushing media should be 40-60 psig air or water depending on customer’s preference.

Signal Processor Control Box

Control box to support two sensors, which can be either/or dissolved oxygen, suspended solids,  pH, ORP, or open channel flow sensors and generate 2 independent and isolated 4-20 mA output signals.  Control box to be designed for future upgrade to commonly used fieldbus protocols by installing a protocol board.  Microprocessor based control box with self-instructing menu, digital circuitry, and illuminated graphical LCD display.  Graphical display to show PH or Redox.    Alternate screen to show the date when the PH sensor was last calibrated (Redox does not require calibration).  The control box produces two linearized 4-20 mA, max 500 ohms galvanically isolated, output signals (12 bit resolution) proportional to pH or ORP.  Control box to have field adjustable contrast from front keyboard panel.  Light on front panel to indicate status of alarm.  To avoid fluctuating output signal, the control box should have a dampening feature or integration time, which can be set from 1 to 999 seconds.  The enclosure should be watertight molded polycarbonate box, NEMA 4X (IP65).  The power supply must be equipped with filter, fuse and varistors for protection against power surges.  The control box should have EEPROM memory, so that following a power failure the control box should start up and resume measuring without requiring recalibration.  Power supply is to be 110/1/60 or 220/1/50.  Control box to be equipped with built-in heater on circuit board to maintain proper temperature inside control box down to –20 F outside temperature.

Programming Module

All programming and settings are performed from the outside of the transmitter by using a self-instructing menu controlled by three touch pad keys.  Special plug-in proms or manually adjustable potentiometers for programming are not acceptable.  Tamperproof programming feature is required to keep settings from being changed, except by authorized personnel.  In the event of power loss, an EEPROM memory should save programming during power outages rather than a battery backup.

Zero point calibration should be made by short circuiting the output of the BNC connector on the amplifier.  The transmitter should have  “smart” and manual calibration features.  “Smart” calibration means that the unit automatically recognizes the pH value of three different buffer solutions.  Buffer solutions for “smart” calibration should be pH values of 4,7 or 10.  When pH calibration is complete, then condition of electrode is to be presented on display as a slope expressed in %.  When performing a manual calibration, pH values of buffer solutions should be entered manually.  When measuring Redox, no other calibration other than zero point should be required.

Flushing program to include independent cleaning function settings for each channel with three (3) programmed functions: 1) intervals from 10 minutes to 99,999 minutes, 2) length of cleaning between 1-99,999 seconds and 3) recovery after cleaning  from 1 to 99,999 seconds.  Output signal should be frozen during cleaning and recovery cycles at initial value so connected equipment will not receive a false pH or Redox readings.

Self-Diagnostics

The software should be of Multi-task design.  It should also contain a watch-dog function connected to the microprocessor.  The software should inform the “watch-dog” at least once per second that the device is working properly.  If it does not, then the “watch-dog” shall restart the processor in order for the unit to resume measuring.

Mounting

The electrode and amplifier to be installed in 1” ( 316SS rod housing by 39” long.  Mounting options are:  1) Standard - SS slide rail assembly mounted to wall for quick and easy removal of sensor or 2) sensor rod installed in 5’ – 10’ telescoping fiberglass rod and supplied with SS spring loaded mounting bracket for 1 ½” handrail.  Options: 1) mounting plate/sun shield from 1/8” aluminum with rain lip complete SS bolts for mounting control box to handrail and 2) flushing nozzle with 33’ hose 1/4” hose and brass solenoid valve, should be available.

Warranty

The manufacturer shall warrant the equipment to be free of defects from workmanship and material for a period of one (1) year after shipment.  Warranty for electrodes to be six (6) months.

Supplier

Specification  is  written  around  equipment  manufactured  by  Cerlic  Controls,  AB,         Model – PHX/REX/BB2,   P/N 25531-49.     Phone  404-256-3097  or  Fax  404-256-3094

File: SPEC/pHX.DOC                                                                                               10/03    rev. 2